## Slide 1: SXR tomography on AUG
Tomáš Odstrčil
todstrci@ipp.mpg.de

## Slide 2: Installation
Not necessary, can be directly started on Linux clusters by command pytomo
Installation from IPP gitlab
module load git
git clone http://gitlab.aug.ipp.mpg.de/todstrci/pytomo.git
cd ./pytomo
./pytomo.sh

## Slide 3: Basic use
1) Select diagnostic
2) Select discharge and time range
3) Check data, remove corrupted LOS, downsample
4) Press start

## Slide 4: Data selection
Check very carefully for corrupted or very noisy channels and remove them
Select one time point to visualize 1D profile and check very carefully for jumps in signal and suspicious values close to zero
In future, also possibility for FFT analysis will be included
1) Moving average, 2) decimate signal
List of wrong LOS

## Slide 5: Reconstruction options
Smoothing
  Isotropic MFI
  Anisotropic MFI – best for SXR
  Isotropic MNGR – very peaked or off-axis SXR profiles
  Anisotropic IMGES – best for divertor reconstruction
Resolution – up to 200x300, time scales linearly with num. pixels
Allow boundary – press reconstruction to zero at boundary (turn off for bolometers)
Anisotropic ration – ratio between parallel and radial smoothing
Scale error bars
  Scale error bars by a constant factor
  Do not affect reconstruction with GCV and PRESS method

## Slide 6: Output panel
Enable output
  Save high quality plots (i.e. pdf)
  Save also all other results to Output path
Separate plots
  Plot emissivity and brightness for each time point
  Can be slow!!
Remove Background – subtract average of all reconstruction or for 10% of time evolution (details in config file)
Plot in loglin scale
  Nonlinear transformation by arcsinh(x/x0)
  For visualizing pedestal region and discharge with a large emissivity range

## Slide 7: Actions panel
SVD
  Remove 1. SVD component from the reconstructions
  Perform SVD analysis
Post processing
  Details about convergence, plasma position, emissivity
Asymmetry
  Calculate poloidal asymmetry
  Compared with simplifies centrifugal model
Sawtooths
  Identify sawtooths times
  Frequency and inversion radius
  Run the tomography for SXR_slow diag without downsampling
Poloidal mode number
  Estimate position and M-number of all modes in the reconstruction
  Use with caution
Impurities
  Calculate W density
  Use IDA if available, else Thomson scattering

## Slide 8: Solvers panel
Ratio presolver
  Estimate and store cross-calibration between cameras
  Run at least ones for each discharge!
Presolvers
  Obsolete option to speed up intermediate steps in MFI regularization
Solvers
  Tiknonov … - original solver, chi2=1 condition, slow O(n3)
  SVD, SVD2, QR, GEV, GSVD – various decomposition methods, SVD2 is the fastest
Lambda solvers
  Select proper regularization
  GCV – preferred, PRESS – more conservative and table, AICc depends on error scale
  Manual – select optimal level by hand – when PRESS and GCV failure
Number of blocks
  speed up reconstruction by solving MFI for block is timeframes
Positivity
  Number gives a weak positivity constrain, button force positivity strongly (slow, can be unstable)

## Slide 9: Transformation panel
Constrain reconstruction on the subspace given by the various radial basis functions
  Abel – pyramid basis
  Cormack - Zernike polynomial
  Fourier-Bessel – Bessel functions
Radial order
  Number of radial basic vectors
Angular order
    Number of angularbasic vectors (cos and sin)
Divertor
  Some attempt to improve smoothing in the divertor region

## Slide 10: Outputs
Emissivity_1.5-5.1_rec_31595.npz
  gres: 3D array, emissivity profile, float16! must by renormalized by gres_norm
  BdMat: average boundary matrix
  Input_parameters – used settings
  Grid points centrods
postprocessing_31595.npz
And many other files… 
Everything stored in ~/tomography/tmp/
When high quality output is enabled also to ~/tomography/Output/
Possibility to write shotfiles, but not integrated in GUI…
Possibility to write ASCII files, but not integrated in GUI…

## Slide 11: Example: SXR reconstruction
Hollow-peaked-hollow profile
Measured brightness mapped on the radial coordinate
Direct comparison between features in data and tomography!

## Slide 12: Example: Bolometers
Special setting necessary! use this command options
pytomo -t 11 -S 7 --positive_constrain -B 200 -u 30 -d 30 --boundary 0 --lambda_solver 0.4 --postprocessing -x 100 -y 150 -s -G 34154
It can be set as default option in the configuration file

## Slide 13: Example: 1/1, 2/2? and 3/2 mode
Reconstruction 29222 at 1.9046s
SVD analysis
Mode numbers profile
M = 1 , 2 ,3

## Slide 14: Example: Impurity analysis
Overestimated
errorbars!
Overestimated
errorbars!

## Slide 15: Example: Poloidal asymmetries
In-out and up-down asymmetry
Simple model for centrifugal asymmetry
If I will find a time, I will implement a new version...

## Slide 16: Example: Sawtooths
Discharge 32595 - current scan from 0.6-1MA
More than 95% crashes detected
High fidelity of inversion radius
Limitations – hollow profiles are difficult

## Slide 17: For experts: config file
Default value for very option in GUI
For example
  Shotfile with mag. equilibrium data
  Generator of phantom profiles
  Radial coordinate rhopol, r/a, rV
  Parameters for sawtooths detector
  Background subtraction options
Plotting options
  Rescale emissivity for each plot separately
  Show LOS, magnetic surfaces, theta star
  Use contours for emissivity, color map
  Plot brightness as function of rho tangential
  SVD chromos in frequency domain
  Number of SVD components

## Slide 18: For experts: command line
Every option in GUI has also command line alternative
pytomo --help for complete list of options
Example
  pytomo 30579 --tmin 4 --tmax 5 -s -u 5 -d 5
  run w/o gui from 4 to 5s smooth and downsample over 5 frames, plot separate plots
Batch processing of discharges is possible
  Missing robust and reliable detection of corrupted channels
  Sometimes optimal regularization is not found

## Slide 19: Geometry calibration
Accurate knowledge of the geometry of SXR diagnostic is essencial for reconstructions
Accuracy of in-situ measurements is usually insufficient
Therefore a self-consistent calibation is necessary and must by performed after every opening or even more often of a systematical discrepancies in backprojected brightness occures
1) select a set of dicharges from peaked to flat without poloidal asymmetry!
2) set this dicharges, times and other parameters in the script detector_optimization.py
3) run the script and wait until the optimization is finished
4) check results, save correction factors and repeat of other discharge

## Slide 20: References
Odstrčil, T., et al. "Optimized tomography methods for plasma emissivity reconstruction at the ASDEX Upgrade tokamak." Review of Scientific Instruments 87.12 (2016): 123505.
Odstrčil, M., et al. "Modern numerical methods for plasma tomography optimisation.„ Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment 686 (2012): 156-161.